**ИНФОРМАЦИОННО СЪОБЩЕНИЕ**

 „АЕЦ Козлодуй” ЕАД уведомява всички заинтересовани лица, че във връзка с постъпил въпрос по пазарна консултация №45778 с предмет „Доставка на Array сонда за контрол на тръбните отвори в колекторите на парогенератори ПГВ 1000 М на 5 и 6 блок в АЕЦ Козлодуй”, прави следните уточнения:

1. Не, ние нямаме array сонда, производство на Zetec.

Вътрешен диаметър на тръбата е 13мм (D=16мм, δ=1.5мм). Забележка: Електрическите характеристики на сондата, трябва да бъдат изчислени за материала на колектора (10ГН2МФА – Руски стандарт).

2.      Манипулатор SGIS е подобен на SM-14, с двоен пушер. Този манипулатор е оборудван с пушер различен от пушер 10D, движещите ролки са размер 5/16 - тази информация е за диаметъра на основното тяло на сондата.

3.      Необходим е само калибрационен блок за X-probe. Прикачената информация е коректна. Мултиплексор не е необходим при условие, че вашите array сонди имат вграден такъв в самата сонда. Slipring: Zetec 36-Pin. Удължителни кабели: ZEC-EXT-PRB, MULTI PURPOSE, 50FT - 3pcs.

**information message**

Kozloduy NPP EAD notifies all interested parties that in connection with a question on market consultation №45778 with subject " Supply of Array Probe for Inspection of the Tube Openings in the ПГВ 1000 М Steam Generator Headers of Kozloduy NPP Units 5 and 6", makes the following clarifications:

1.      No, we don’t have Zetec array probe. Tube ID is 13mm (OD=16mm, wall thickness= 1.5mm). NOTE: Electrical characteristics of the probe should be calculated for tube sheet material (10ГН2МФА Russian standard).

2.      Manipulator SGIS is some like SM-14, with dual probe push-puller. This robot is equipped with pusher other than 10D, the drive wheels are 5/16 - this information is for the diameter of body probe.

3.      We need just X-Probe calibration standard. The attached information is correct. Multiplexer is not necessary in case, that your array probes have built-in within the head of the probe. Slipring: Zetec 36-Pin. Extension cables: ZEC-EXT-PRB, MULTI PURPOSE, 50FT - 3pcs.